Objective

To define the procedure for the control of inspection, measurement, and test equipment. This procedure will also ensure that the equipment is capable of the required accuracy and is properly calibrated.

Note: For equipment that does not require calibration but needs systematic validation, see QMS work instruction OPS-W-835-01 Test Equipment and No Calibration Required Control.

This procedure is the corporate level document that is in place when site level processes are not approved by a regulatory authority or directed by customer requirement. This procedure shall be followed when no other approved process is in place and is the default for all sites.

This procedure applies to inspection tools, measuring devices, and all other test and measurement equipment used by employees of Burrana for determining conformity of products and services to requirements.

Responsibilities

This procedure is intended for all Burrana departments where measurement is needed to assure conformity of product.

Utah, Calexico, and Toulouse are domestic centers with regulatory oversight, if they need additional changes outside of the actions established in this procedure, they must inform the appropriate regulator of any changes made, taking into consideration the relevant Airworthiness Authority regulations for that facility.

Definitions & Acronyms

Definitions

| TERM | MEANING |
| --- | --- |
| Burrana | Any organization in the Burrana Group of companies. |
| Calibration Coordinator | The person appointed to control calibration within a location, “Company Calibration Coordinator” is the calibration Coordinator for all Burrana. |
| Gauge or Gage or IMTE | Refers to any inspection, measuring and test equipment used to verify the conformity of products and services to the requirements. |
| QAS | Quality Assurance Support personnel under the direction of the Global Manager of Quality Assurance. |
| Verification | Internal process to review tools using certified standards for the purpose of certifying those tools for use in production or repair. |
| Validation | Calibration through a 3rd party that has been accredited to ISO17025 and carries the NATA/NIST approvals |

Acronyms

| TERM | MEANING |
| --- | --- |
| A2LA | American Association for Laboratory Accreditation. (www.a2la.org) (US) |
| AU | Australia. |
| GR&R | Gage Repeatability and Reproducibility. |
| NATA | National Association of Testing Authorities. (AU) |
| NIST | National Institute of Standards and Technology. (US) |
| US | United States. (of America) |
| VP | Vice President. |

Process

General

The responsibility for ensuring compliance with this procedure is with the Manager of each department with support from Quality Functions.

It is the responsibility of all employees to verify the accuracy of calibrated tooling before use in a process.

The responsibility for the safeguard of the gauge is the user. The user shall ensure proper handling, usage, and storage of the gauge.

Gauges must be calibrated, certified, and labelled before use. Any gauge or tool found to be overdue for calibration shall be immediately surrendered to Calibration Coordinator for a change to inactivate status and a non-conformance report issued. The exception is that a non-conformance report is not required if that gauge is from the Engineering department and has not been in use during the un-calibrated period.

Any personal gauges used at Burrana may be entered into the calibration system and must be kept current. Personal gauges may be marked “Reference Only” and may NOT be used in any way for inspection data, design data collection/determinations, or to determine product conformity.

Calibration of equipment and gauges shall be performed by an approved calibration laboratory with standards traceable to National Institute of Standards and Technology (NIST) or other accrediting national standards authority (i.e., NATA in Australia) and preferably A2LA (US) accredited. Documentation must reflect which standards they are traceable to.

Unless otherwise stated, v-blocks, and mounting fixtures shall be treated as “for indication only”

and are not subject to calibration.

New Gauges

Prior to the use of any new gauge for the first time, the gauge shall be brought to the Company Calibration Coordinator for entry and recording into the calibration system by following the Equipment Receiving Flow Chart from Appendix 1.

The Calibration Coordinator shall assign the new equipment a reference number. This number shall be affixed to the gauge using the most permanent and practical method possible. In cases where it

is not practical to mark the gauge, the storage box may be marked, or the equipment must be tagged. In these cases, the gauge must always be stored in the box, or the tag must always remain with the equipment. The identification should be per QMS procedure OPS-P-715-01 - Equipment Identification.

The Warehouse must inform the Company Calibration Coordinator when calibrated gauges have been received. Only Test Equipment Team members or department managers are to receive Calibrated Equipment.

New gauges shall have a new record entered the Burrana calibration system by the Company Calibration Coordinator before use. The information should contain the following:

Reference ID number

Serial Number (if applicable)

Model number

Gauge type

Division

Storage Location

Manufacturer

Typical unit of measure (distance, weight, angle, etc.)

Calibration Expiry

The local Calibration Coordinator shall review the calibration due report and determine which gauges are due for calibration in the current and the following month. It is strongly recommended that gauges requiring calibration in the current month be removed from service prior to the calibration expiration date.

The gauge shall be segregated in an identified secure area to prevent unintentional use until the gauge can be sent out for calibration.

The Calibration Coordinator will record the following information for gauge traceability as appropriate.

The date that the gauge was sent out for calibration.

The name of the company who will be performing the calibration.

The expected due date of the gauge.

Any known problems or issues with the gauge will be provided to the organization carrying out the calibration.

Gauges returned from Calibration

Upon return of the gauge from calibration, the Calibration Coordinator shall verify that the gauge is labelled with the equipment reference number (or certificate number), the date of calibration, and the due date for the next calibration. The Calibration Coordinator shall review the Calibration Report to confirm the equipment was not greater than 10% out of tolerance prior to adjustment. Any gauge above this level shall comply with section 4.5.3.

The Calibration Coordinator shall also verify that the Calibration Certificate is acceptable. The certificate must include:

Identify the equipment reference number and/or serial number.

Be signed and/or stamped by the calibration laboratory.

Confirm that the gauge is received and returned within the specified tolerance condition.

Provide the traceability to NIST (US), NATA (AU) or equivalent.

Note that no defective condition is noted or reported by the calibration laboratory.

Indicates that a complete calibration was performed on the gauge. For example, calipers require OD, ID, and depth calibration.

Show the “As received” and the “As returned” conditions of the gauge.

Out of Tolerance Measuring Equipment and Tools

Gauges which are found to be out of tolerance or not properly calibrated, shall be tagged and segregated in a secure area per section 4.3.2 by the Calibration Coordinator to prevent its use.

The gauge shall be tagged “DO NOT USE” and should include the failure reason and the actual date.

The Calibration Coordinator shall review the reported gauge condition and must create a SOOT analysis. The Calibration Coordinator shall create the SOOT with the information provided as follows:

Advise the Supervisor of the area and the Quality Engineer of the out of tolerance.

The Supervisor is to confirm if the gauge was used in the rejected range for any validation measurement and advise the Calibration Coordinator of the result.

Report to the Quality Manager to determine if a product recall & notification to customers is necessary and/or required.

Calibration Intervals

Tools requiring calibration are periodically calibrated in accordance with the tool manufacturers’ published standards and recommendations. Where no recommendation for the calibration is published (such as, but not limited to ANSCI/NCSLZ540-11994, ANS/ISO/IEC 17025:2005, or ANSI/NCSL Z540.3-2006) or where the calibration methods or standards are not specified, calibration is carried out in accordance with the requirements of the ISO 10012. These standard details both the generic requirements and guidance for the implementation of measurement management systems.

Specific recalibration intervals are dependent on several factors including:

Manufacturer’s calibration recommendations.

Previous calibration history.

Accuracy requirements set by customers.

Requirements set by contract or regulation.

Inherent stability of the specific instrument or device.

Environmental factors that may affect stability.

Frequency of use.

Internal Verification of Tools

Only an approved technician can identify tools to be verified vs calibrated.

The calibrations technician is approved by the definition of the position.

Where a tool has been identified to be verified internally, Burrana will record the data for the standard used to verify the tool and create a certificate in association with the verification certificate.

If during an internal verification a tool fails to test within the standard requirements, that tool will be sent for calibration and an investigation opened in the data control system using the NC module. The process QC-P-870-01 Control of Nonconforming Outputs will be implemented, and the results recorded.

Verification does not exclude the recall process and will follow those same requirements for recall as described in the recall section of this document.

Adjustment of Calibration Interval

For tools such as Tape Measurers, Steel Rules, or other stable measurement devices, the organization has determined that the RISK to product non-conformity is low to insignificant when the   cycle   is   extended   to   3    years    including    the    first    calibration    cycle    expiration. This cycle is to be adjusted in the system but requires that the technicians review the tools for defects or damage prior to use. If any damage is identified, the tool is immediately segregated and tagged for disposal. This process is not considered to be an extension, but a standard calibration cycle for static measuring devices and is excluded from the requirement set forth in this procedure.

The Calibration Coordinator or Quality Engineer may recommend to the Quality Manager that a gauge can have its calibration cycle extended, based on criteria described in section 4.6. In special cases an extension may be granted to complete a current job where there is no suitable replacement gauge.

A basic Risk assessment must be completed under the review tab in the control system software to show justification for the extension of cycle for any item of calibrated equipment.

If the gauge has a history of three (3) or more successive calibrations where no adjustments or repairs are made, then the calibration cycle may be lengthened at the discretion of the Calibration Coordinator or Quality Engineer but not exceed manufacturers specifications. Note: This does not apply for Service Center´s equipment.

Where an extension is approved, the extension time shall not exceed six (6) months from the date the calibration was due. The Calibration Coordinator will document this extension by putting an extension sticker or tag on the gauge. This sticker will include the words ‘Extended Calibration’, the gauge reference number, and the extended calibration due date per QMS procedure OPS-P-715-01

- Equipment Identification.

If the gauge needs readjustments or repairs twice consecutively, then the calibration interval may be shortened to one-half of the present calibration cycle or other interval as approved by the Calibration Coordinator or Quality Engineer.

Damaged or Inaccurate Gauges

If any gauge is dropped or discovered to be damaged while in service, that gauge shall be surrendered to Calibration Coordinator immediately.

The Calibration Coordinator shall review and disposition the gauge per QMS procedure CORP-P- 713-03 IT-Test Equipment Disposition.

Calibration Coordinator may arrange for the repair and re-calibration of the gauge if practical.

Qualification of Calibration Laboratories

Burrana does not, generally, perform any in-house calibration. Exceptions may be for specialty mounting fixtures where specific dimensions of the mount are required. Refer QMS procedure CORP-P-713-04 – Maintenance Control.

Laboratories used for Calibration shall be:

Approved per the requirements of the applicable Burrana SOP for Supplier Approval & Evaluation.

Meet the requirements of ISO10012, or ANSI/NCSL Z540-1 and maintain traceability of all measurement standards to the NIST, or equivalent standard (i.e., NATA).

Preferred accreditation through A2LA (American Association for Laboratory Accreditation).

Calibration Labs are required to include the standard used to perform the calibration including the P/N, S/N, and Calibration due date on all certificates.

Test equipment and gages used in Engineering

Within Burrana Engineering workshops, some gauges are not used for determining conformity of products and services to requirements, and tooling which is only required for specific projects. To provide a more efficient system, such equipment may be tagged as inactive within the data control system when short-term or long-term calibration is no longer required.

As Engineering IMTE becomes due for calibration, the area manager will identify the need to either inactivate or re-calibrate IMTE based on its forecast use in determining product conformity and validation. If calibrated IMTE is to become inactivated, the equipment is to be clearly marked as uncalibrated, and any previously expired calibration sticker must be removed or covered. Refer to Project Specific Calibration ENG-W-715-01.

Test equipment and gages used in Operations

Daily Check – Care and Use of Equipment

To assure calibrated tools and gauges used for product or article acceptance are protected, maintained, and used in an acceptable environment, they shall be secured in their designated areas.

Portable equipment shall be secured in a designated location and issued for use as required using an equipment checkout log.

Before starting the inspection or test of any part:

Check all equipment for any visible damage. Report any damage or excessive wear to TEE.

Check the ZERO setting if apply. Ensure the instrument is properly zeroed (or zeroed-out).

Check the equipment against a standard gauge block (or other similar standard) to ensure a proper reading. For example, it is good practice to check calipers or micrometers using a set of calibrated gauge blocks to ensure the accuracy of the calipers/micrometers.

For internal control, when equipment and work tools assigned to an area are required, it must be

registered in the QMS form ‘OPS-F-810-10 - Tool Control’ to log delivered and returned tooling.

Environmental Conditions

As applicable, environmental conditions for inspections and measurements will be controlled to a degree to ensure accuracy and repeatability of measurements. Room temperature will be at a comfortable level and temperature control will become necessary in extreme hot and cold weather environments. Equipment relocated from different temperatures will be allowed to stabilize before use.

In situations where temperature and humidity are factors, regular checks shall be performed to verify the appropriate environmental conditions. In such cases, if the environmental conditions are outside of the prescribed range, inspections and/or measurements should be suspended until the environmental conditions have been controlled NOTE: The Calibration Coordinator may authorize an inspection or measurement of a part if it is determined that the environmental condition will not affect the conformity of the inspected feature or part. For example, humidity will not have as significant an effect on measurements of electronic components as it would for elastomeric parts.

Measurement System Analysis (MSA) and Gage R&R

The purpose of this section is to determine if Burrana’s measurement system is capable to generate accurate data and to determine if the accuracy of the data is adequate to achieve the Operations and Quality goals.

The main goal is to also define if the differences of the data acquired are due to actual differences in what is being measured and NOT to VARIATION in measurement methods.

Only applies for measurement tools, Automated Test Equipment is not applicable for GR&R.

Criteria for GR&R study:

o The equipment that needs to be analyzed with GR&R will be defined in the data control system with a frequency of 1 year and the result will be stored in their equipment´s profile as well.

GAGE R&R

This analysis helps us to demonstrate the ability to get the same result if we take the same measurement repeatedly or if different people take the same measurement, the GR&R involves evaluating REPEATABILITY and REPRODUCIBILITY of a measurement system.

REPEATABILITY

Refers to the variability of the measurement system and it is the variation that occurs when successive measurements are made under same conditions: SAME PERSON, same object measured, same characteristic, same instrument, same setup.

REPRODUCIBILITY

Is the variation in the average of measurements made by different operators using the same measuring instrument and technique when measuring the identical characteristic on the same part and same process: DIFFERENT PERSON, same object measured, same characteristic, same instrument, same setup.

Governing Documents

	AS9100

	AS9110

	AS9115

	AS9145

	AS9120

Forms and Documented Information

Forms

	Calibration Certificates and Records.

   Repair Station Equipment List.

 The authorized Certificates of Calibration provide evidence that the that the equipment and gauges have been calibrated in accordance with this procedure. The records of Calibration (hard copy) shall be kept and stored in accordance with applicable QMS procedure QA-P-750-01 - Control of Documented Information.

	Calibration records (electronic files) for each gauge, tool, or equipment shall be retained and stored such that it is readily accessible to QAS.

Documented Information

	CORP-P-600-01 – Planning for the Quality Management System.

	CORP-P-810-01 – Operational Planning and Control.

	CORP-P-713-03 – IT-Test Equipment Disposition.

	CORP-P-713-04 – Maintenance Control.

	PUR-P-840-01 – Control of Externally Provided Processes, Products and Services.

	QC-P-715-02 – Equipment Identification.

	QC-P-851-04 – Control of Production and Service Provision.

	QC-P-870-01 – Control of Nonconforming Outputs.

	SCI-P-910-01 – Monitoring, Measurement, Analysis and Evaluation.

	QA-P-750-01 – Control of Documented Information.

	OPS-P-835-01 – Test Equipment and No Calibration Required Control.

	QC-F-851-02 – GR&R Data Form.

	OPS-F-810-10 – Tool Control.

	F-715-02 – Calibration Inspection.

Opportunities and Risks

The planning procedure CORP-P-600-01 for Planning for the Quality Management System addresses opportunities and risks / (risk-based thinking).

Burrana Makes use of organizational knowledge, lessons learned and experience with the activities associated with Control of Monitoring and Measuring Equipment to determine opportunities and risks that need to be addressed, and which can:

Give assurance that the procedure can achieve its intended results.

Enhance desirable effects and prevent or reduce undesired effects.

Achieve improvement. Whenever necessary this information is documented.

Appendix 1

Table 1 - Equipment Receiving Flow Chart

Upload into QwizPro

Is NCR

equipment

No

Yes

equipment,

Test equipment,

Test equipment , Need to Assignee a label with a Burrana ID, Also NCR ID label if is required

Yes

Is a

Purchased Tool?

No

No

Need a Destructive tests or set up from production Engineering

Yes

Upload Test equipment

data base

Production Engineering need to provide certification Method

PO need to be closed by Warehouse Personnel then delivered only for test equipment Coordinator

All transferred Tool or equipment must be delivered only to Test equipment Coordinator

Receiving Tool or equipment from Warehouse

Upload into QwizPro

Is NCR

equipment

No

Yes

equipment,

Test equipment,

Test equipment , Need to Assignee a label with a Burrana ID, Also NCR ID label if is required