# Purpose

- To define the procedure for the control of inspection, measurement, and test equipment. This procedure will also ensure that equipment is capable of the required accuracy and is calibrated.

- This procedure applies to inspection tools, measuring devices, and all other test and measurement equipment used by employees of Techno-Aide for acceptance of product or design of product and distribution.

Responsibilities

- This procedure is intended for Techno-Aide Receiving Inspection and Quality Assurance departments.

Definitions

| Term | Definition |
| --- | --- |
| A2LA | American Association for Laboratory Accreditation (www.a2la.org) (US) |
| The Organization | Techno-Aide name |
| Gauge or Gage | Refers to any inspection, measuring and test equipment |
| NIST | National Institute of Standards and Technology (US) |
| QUALITY MANAGER | Quality Assurance Support personnel under the direction of the Quality Manager |
| US | United States (of America) |

Equipment or Software

- M2M

Process

- The responsibility for compliance with this procedure is with the Production Manager

- Responsibility for the safeguard of the gauge is the user. The user shall ensure proper handling, usage, and storage of the gauge.

- Gauges must be calibrated, certified, and labeled before use. Any gauge or tool found to be overdue for calibration shall be immediately surrendered to the Production Manager for quarantine.

- Any personal gauges used at Techno-Aide may be entered into the calibration system and must be kept current. Personal gauges may be marked “Reference Only” and may NOT be used in any way for inspection data, design data collection/ determinations, or to determine product conformity.

- Calibration of equipment and gauges shall be performed by an approved calibration laboratory with standards traceable to National Institute of Standards and Technology (NIST) or other accrediting national standards authority and preferably A2LA (US) accredited.

- Unless otherwise stated, v-blocks, and mounting fixtures shall be treated as “for indication only” and are not subject to calibration.

New Gauges

- Prior to the use of any new gauge for the first time, the gauge shall be brought to the PRODUCTION MANAGER for entry and recording into the Calibration system.

- PRODUCTION MANAGER shall assign the new gauge a gauge number. This number shall be affixed to the gauge using the most permanent and practical method possible. In cases where it is not practical to mark the gauge, the case may be marked, or the gauge must be tagged. The case or the tag must always remain with the gauge.

- The new gauge shall have a new record entered in the Techno-Aide calibration system. The information should contain the following:

  - Gauge ID number

  - Gauge Serial Number (if applicable)

  - Model number

  - Gauge type

  - Storage Location

  - Manufacturer

  - Typical unit of measure (distance, weight, angle, etc.)

Gauge Calibration Recall Procedure

- PRODUCTION MANAGER shall review the calibration due report and determine which gauges are due for calibration in the current and following month. It is strongly recommended that gauges requiring calibration in the current month be removed from service prior to the calibration expiration date.

- The gauge shall be segregated into a secure area to prevent the unintended use until the gauge can be sent out for calibration.

- PRODUCTION MANAGER will record the following information for gauge traceability as appropriate.

  - The date that the gauge was sent out for calibration.

  - The name of the company who will be performing the calibration.

  - The expected due date of the gauge.

  - Any known problems or issues with the gauge.

Gauges received from the Calibration Lab

- Upon receipt of the gauge from the calibration lab, PRODUCTION MANAGER shall verify that the gauge is labeled with the correct serial number, gauge ID (or certificate number), the date of calibration, and the due date for the next calibration.

- PRODUCTION MANAGER shall also verify that the Certification of Calibration is acceptable. For example, the Certification:

  - Identify the gauge ID and/or serial number.

  - Is signed and/or stamped by the calibration laboratory.

  - Confirms that the gauge was received and returned within the specified tolerance condition.

  - Provides the traceability to NIST (US) or equivalent.

  - That no defective condition was noted or reported by the calibration lab.

  - That a complete calibration was performed on the gauge. For example, calipers require OD, ID, and a depth calibration.

  - Must show the “as received” and the “as returned” conditions of the gauge.

Out of Tolerance Gauges

- Gauges which are found to be out of tolerance, damaged, broken, or not properly calibrated, the PRODUCTION Manager or PRODUCTION MANAGER shall Tag and segregate the gauge in a secure area to prevent its use.

- The Tag shall state “DO NOT USE” and should include the reason, such as “out of calibration” or “broken” and a date.

- The PRODUCTION Manager shall review the reported gauge condition and determine what action is required. In the event any gauge is found to be out of tolerance by more than 10% of the accuracy of the gauge, the PRODUCTION Manager shall:

  - Place on “QC HOLD” all products which have been inspected with the gage during the time in which the gauge was suspected of being out of tolerance.

  - Determine if a product recall & notification to customers is necessary and/or required.

NOTE: 10% accuracy refers to the least significant value (or smallest graduation) of the gauge. Product affected may be tracked using Inspection Records, Batch Numbers (BN), Sales History, et al.

Calibration Intervals

- Test and measurement equipment shall be calibrated at prescribed intervals. Calibration intervals shall be based upon manufacturer’s recommendations, previous calibration history, and the stability of the gauge. The following gauge family/type intervals are to be used as a guideline only:

| Tool, Gauge, or Equipment | Recommended Calibration Interval |
| --- | --- |
| Precision hand tools, micrometers, calipers, dial indicators, linear height gauge, and other hand tools Torque Wrenches & Screwdrivers Oscilloscopes, Digital Multi-Meters, Test Set, High Voltage Insulation Testers, etc. | 12 months |

Adjustment of Calibration Interval

- The PRODUCTION Manager may determine that a gauge can have its calibration cycle extended, based on criteria such as manufacturer calibration frequency suggestions, frequency of use, or evidence of continued stability. In exceptional cases an extension may be granted to complete a current job where there is no suitable replacement gauge.

  - If the gauge has a history of five or more successive calibrations where no adjustments or repairs are made, then the interval may be lengthened at the discretion of the PRODUCTION Manager or Quality Engineer.

  - If the gauge needs readjustments or repairs twice consecutively, then the calibration interval may be shortened to one-half of the present calibration cycle or other interval as approved by the PRODUCTION Manager or Quality Engineer.

- Where an extension is approved, the extension time shall not exceed six (6) months from the date the calibration was due. PRODUCTION MANAGER will document this extension by putting an extension sticker or tag on the gauge. This sticker will include the words “Extended Calibration,” the gauge number, and the extended calibration due date.

- If additional time is needed after the six (6) month period, only the PRODUCTION Manager may extend the calibration due date for another three (3) months. Under no circumstance can the gauge be extended past nine (9) months from the original calibration due date.

Damaged or Inaccurate Gauges

- If any gauge is dropped or discovered to be damaged or inaccurate while in service, that gauge shall be surrendered to the PRODUCTION MANAGER immediately.

- The PRODUCTION Manager shall review and disposition the gauge per this procedure.

- PRODUCTION MANAGER may arrange for the repair and re-calibration of the gauge if practical.

Qualification of Calibration Laboratories

- Techno-Aide does not, generally, perform any in-house calibration. Exceptions may be for specialty mounting fixtures where specific dimensions of the mount are required.

- Outside Calibration Laboratories shall be:

  - Approved per the requirements of the applicable Techno-Aide SOP for Supplier Approval & Evaluation.

  - Meet the requirements of ISO17025, or ANSI/NCSL Z540-1 and maintain traceability of all measurement standards to the NIST, or equivalent standard (i.e., NATA).

  - Preferred accreditation through A2LA (American Association for Laboratory Accreditation – www.a2la.org).

Daily Check – Care and Use of Equipment

- To assure calibrated tools and gauges used for product or article acceptance are protected, maintained, and used in an acceptable environment, they shall be secured in their designated areas. Portable equipment shall be secured in a designated location and issued out for use as required using an equipment checkout log.

- Before starting the inspection of any part:

  - Check all gauges for any visible damage. Report any damage or excessive wear to QUALITY MANAGER.

  - Check the ZERO setting. Ensure the gauge is properly zeroed (or zeroed out).

  - Check the gauge against a standard gauge block (or other similar standard) to ensure a proper reading.

    - For example, it’s good practice to check calipers or micrometers using a set of calibrated gauge blocks to ensure the accuracy of the calipers/micrometers.

Environmental Conditions

- As applicable, environmental conditions for calibrations, inspections, and measurements will be controlled to a degree to ensure accuracy and repeatability of measurements. As required, inspections and measurements shall occur under standard laboratory conditions

- In situations where temperature and humidity are factors, regular checks shall be performed to verify the appropriate environmental conditions. In such cases, if the environmental conditions are outside of the prescribed range, inspections and/or measurements should be suspended until the environmental conditions have been controlled.

Control of Revision Table

| Revision | Changes | Created by | Approved by | Date of approval |
| --- | --- | --- | --- | --- |
| Draft | New document | Brian Olea (OAQC) |  |  |