# Objective

This procedure ensures that all monitoring and measuring equipment (MME) used to verify product or process conformity is suitably selected, calibrated or verified, safeguarded against unauthorized adjustment or damage, and that any out-of-calibration condition triggers impact investigation, corrective action, and customer notification, fully meeting ISO 9001:2015 Clause 7.1.5 and related requirements.

# Applicability

This procedure applies to every device, gauge, fixture, software tool, and reference standard whose quantitative measurements or pass/fail outputs provide evidence of conformity during design verification, production operations, inspection activities, or servicing of industry, space, or defence products within the scope of the QMS.

# Responsibilities and Authorities

The Quality Manager approves this procedure and reports MME performance metrics—including on-time calibration rates and out-of-calibration events—in each Management Review.
The Equipment Coordinator verifies equipment suitability prior to procurement, enters all MME into the Calibration Register (F-715-002), schedules and conducts calibrations or verifications, affixes standardized calibration labels, seals adjustment points, and quarantines nonconforming items.
Department Managers ensure that only equipment bearing current calibration labels is used, verify label currency before use, and enforce compliance with WI-715-002.
The IT Manager secures electronic calibration records and software-validation logs with access controls and nightly backups in accordance with P-712-01.
External Calibration Laboratories deliver ISO/IEC 17025-accredited services and provide traceable certificates; their selection and performance are managed under P-840-01.
All employees must use only labeled equipment, handle it per prescribed preservation requirements, and report any anomalies via the Equipment Problem Report (F-710-001).

# Definitions

| Term | Definition |
| --- | --- |
| Monitoring and Measuring Equipment (MME) | Any device, gauge, fixture, software application, or reference standard that produces quantitative measurements or pass/fail results used as evidence of product or process conformity. |
| Calibration Register | The controlled record (F-715-002) listing each MME item’s unique ID, model, serial number, tolerance, intended use, calibration interval, last calibration date, next due date, and owner. |
| Calibration Label | The standardized label displaying equipment ID, calibration date, next due date, and coordinator initials that must be affixed to each calibrated or verified item. |
| Key Characteristic | A feature of a product whose variation has a significant effect on fit, form, function, performance, or service life, as defined in design output documentation. |

# Procedure

## Determination of Monitoring-and-Measuring Equipment Requirements

The Quality Team and AS Steering Committee determine the types and quantities of monitoring-and-measuring equipment (MME) required to obtain valid and reliable evidence of conformity, in alignment with Operational Planning and Control (P-810-01), Control of Production and Service Provision (P-851-01), and Monitoring, Measurement, Analysis and Evaluation (P-910-01).

## Initial Verification and Entry to Service

Each new or modified MME item undergoes initial verification. Equipment that passes is released into service; unverified or failed items remain in quarantine until verification is completed and documented.

## Equipment Identification and Register Control

The Equipment Coordinator assigns every MME item a unique identification number and records it in the Calibration Register (F-715-002), capturing equipment name, manufacturer, model, serial number, organizational asset ID, initial verification date, calibration requirement, maintenance requirement, calibration interval, responsible owner, last calibration date, next due date, record-storage location, equipment location, tolerance levels, and specified environmental conditions.

## Calibration and Verification Scheduling

MME items are scheduled for calibration or verification at defined intervals and, where required, prior to critical measurement activities. All activities are performed under environmental conditions that meet or exceed those specified for use. Traceability to national or international standards is maintained for customer-designated Key Characteristics and where contract or regulatory requirements apply.

## Pre-Use Verification

Users verify that each MME item bears a current calibration label before every use. Any item found overdue, unlabeled, or otherwise unfit for purpose must be removed from service and placed in quarantine.

## Employee Notification and Register Maintenance

Employees notify the Equipment Coordinator when MME items must be added to or removed from the Calibration Register, and report any equipment concerns via the Equipment Problem Report (F-710-001). The Equipment Coordinator reviews and updates F-715-002 accordingly.

## Maintenance, Preservation, and Sealing

All maintenance and calibration follow supporting work instructions or manufacturer’s manuals, which specify methods, acceptance criteria, calibration-equipment certification, scope and range of use, and storage or environmental requirements. Adjustment points on MME are sealed to prevent unauthorized tampering, and equipment is handled, preserved, and stored to protect its accuracy and fitness for purpose.

## Quarantine and Impact Investigation

When an MME item is discovered to be overdue, out of tolerance, or otherwise unfit for use, it is immediately placed in quarantine. The Equipment Coordinator conducts an impact investigation within one working day, documenting affected product lots or inspection results in the Out-of-Calibration Investigation Log (tab in F-715-002). Affected product is controlled per P-870-01, systemic causes trigger P-880-01 Corrective Action, and any potentially affected customers are notified in writing within two business days under P-820-01.

# Records Retention and Backup

All calibration certificates, investigation logs, and software-validation records are retained in the controlled document management system for a minimum of five years and are protected by nightly electronic backups as per P-712-01.

# Performance Monitoring and Continual Improvement

Monthly performance metrics—including on-time calibration rate, number of out-of-calibration events, and software-validation compliance—are collected and reviewed during Management Review. Improvement opportunities are recorded on F-610-001 and tracked to closure.

# References

• ISO 9001:2015 Clause 7.1.5
• ISO/IEC 17025:2017 – General requirements for the competence of testing and calibration laboratories
• P-710-01 Resource Management
• P-712-01 Information-Technology & Cybersecurity
• P-820-01 Customer-Related Processes
• P-840-01 Control of External Providers
• P-870-01 Control of Nonconforming Outputs
• P-880-01 Corrective Action
• WI-710-002 Environmental Condition Monitoring
• WI-715-002 Calibration & Verification Work Instruction

# Documented Information

- Calibration Register (F-715-002)

- Equipment Problem Report (F-710-001)

- Out-of-Calibration Investigation Log (tab in F-715-002)

- Calibration Certificates

- Quarantine Records

# Revision History

| Rev | Date | Summary of Change | Authorised By |
| --- | --- | --- | --- |
| A | [Orig] | Initial issue in compliance-only format | — |
| B | [Date] | Added design-development linkage and customer notification | — |
| C | [Date] | Reformatted into section/subsection layout; refined quarantine language | [Approver] |

# Record of Review

| Review Date | Reviewed By | Summary of Changes | MRM Reference |
| --- | --- | --- | --- |
|  | OQCQMS | Rev C issued – sections and subheadings applied | OQCQMS |