# Purpose

Describe how [The Organization] controls, recalls, and records calibration or verification of monitoring-and-measuring equipment (MME) to ensure product conformity meets AS9100 Rev D Clause 7.1.5 requirements.

# Scope

Applies to all MME—physical, mechanical, electrical, or software—that can influence acceptance decisions for a product or process.

# Responsibilities

| Role | Responsibility |
| --- | --- |
| Calibration Coordinator | Maintains F-715-001 register; issues recall notices; reviews out-of-tolerance (OOT) reports. |
| Process Owners / Technicians | Submit equipment for service on or before due date; quarantine OOT product pending disposition. |
| Quality Manager | Approves service providers; analyses OOT impact; reports KPIs at Management Review. |

# Definitions

| Term | Definition |
| --- | --- |
| MME | Monitoring- & Measuring-Equipment used to verify conformity. |
| Calibration | Comparison against a standard of known accuracy to establish traceability. |
| Verification | Confirmation that specified requirements have been met (may not adjust the device). |
| OOT | Out-Of-Tolerance—device found outside specified limits during service. |
| Recall List | Monthly list of devices due for calibration/verification within the next 30 days. |

# Work Instruction

| Step | Action | Responsible | Record |
| --- | --- | --- | --- |
| 1 | Identify & Register MME – On receipt or creation of new equipment, enter details into Measuring-Equipment Register (F-715-002). | Calibration Coordinator | F-715-001 |
| 2 | Assign Unique ID & Label – Affix durable label showing ID, last service date, next due date, and status (“CAL” or “VER”). | Calibration Coordinator | Device label |
| 3 | Schedule Service – Generate monthly Recall List from F-715-002 and email to Process Owners. | Calibration Coordinator | Recall List |
| 4 | Send Device for Calibration/Verification – Before or on due date, log transfer in F-715-002 and ship to approved lab. | Process Owner | F-715-001 |
| 5 | Receive & Review Cert – Verify certificate includes standards traceable to NIST (or national equivalent) and measurement results. | Calibration Coordinator | Certificate |
| 6 | Update Register – Record new service date, due date, and results in F-715-002. File certificate electronically. | Calibration Coordinator | F-715-001 |
| 7 | Handle OOT Results – If device is OOT: ① Tag device “Do Not Use”; ② Raise Out-of-Tolerance Report (F-715-003); ③ Quality Manager performs impact analysis on product accepted since last in-tolerance date. | Calibration Coordinator & Quality Mgr | F-715-001 |
| 8 | Quarantine Suspect Product – Segregate product affected by OOT device and process via P-870-01. | Process Owner | NCR / F-870 forms |
| 9 | Verify Adjustment Marks – After repair/adjustment, confirm device readings at at least two points across range. | Calibration Coordinator | Certificate |
| 10 | Control Software & Firmware – Validate measurement software version; record checksum in F-715-002. | Calibration Coordinator | F-715-001 |
| 11 | Review KPIs – Monthly: % on-time calibrations, OOT rate; present trends at Management Review. | Quality Manager | KPI chart |
| 12 | Retire or Re-purpose – For obsolete or unrepairable MME, update status to “Retired” in F-715-002 and remove from service area. | Calibration Coordinator | F-715-001 |

# Review & Update Frequency

- Planned review: annually.

- Immediate update upon significant process/equipment changes or audit finding.

# References

- AS9100  — Clause 7.1.5

- P-715-01 Control of Monitoring & Measuring Equipment (Procedure)

- P-870-01 Control of Non-conforming Outputs (Procedure)

- P-750-01 Control of Documented Information (Procedure)

# Documented Information

| Identifier | Title | Purpose |
| --- | --- | --- |
| F-715-001 | Measuring-Equipment Register | Master list with status & due dates |

Record of Review

| Review Date | Reviewed By | Summary of Changes | MRM Ref. |
| --- | --- | --- | --- |
| – | – | Initial release | – |